## Slide 1: Electrode Training
HORIBA Instruments, Inc.
Scientific
Pamela Millett
May 15, 2015

## Slide 2: pH Electrodes
What is G, R, T?
  G = Glass (the pH sensor)
  R = Reference sensor
  T = Temperature sensor
    3 in 1: includes all sensors, G,R & T
    Combination Electrodes: include G & R
    Glass Electrodes: pH sensor only (G)
    Reference Electrodes: reference sensor only (R)
    Temperature Electrodes: temperature sensor only (T)

## Slide 3: pH Glass Electrode Method
A pH and reference electrode are required to measure pH
The voltage (potential) is measured between the pH and reference sensors
Most of HORIBA’s electrodes are glass electrode
Glass electrode method is the most common and reaches equilibrium quickly

## Slide 4: HORIBA pH Buffer Setting Options
Options: NIST, USA or Custom (CUST)
What is NIST? National Institute of Standards and Technology
NIST Values: 2=1.679, 4=4.008, 7=6.865, 9=9.180, 12=12.454
USA Values: 2=2.000, 4=4.000, 7=7.000, 10=10.000, 12=12.000
CUST Values: Your choice for up to 5 calibration points

## Slide 5: ATC or MTC Temperature Settings
ATC is the default setting
ATC: Automatic Temperature Compensation
Can only be used if a temperature sensor is
included in the electrode (GRT or GR+T)
MTC: Manual Temperature Compensation
Temperature is input using an external temperature device
Sometimes used when extremely precise measurements are required using a constant temperature bath
Temperature is important for pH measurements

## Slide 6: Semiconductor pH Electrode Method
ISFET: Ion Selective Field Effect Transistor
Sensor does not have glass, popular for food samples
Model 300 Solution: Included for the zero adjustment
Preferred for durability
Protect from light: Sensor is light sensitive
Principle: Current flowing through two electrodes is measured at the “gate” electrode that is in contact with the sample

## Slide 7: Ion Electrodes
Condition the electrodes before use per the instruction manual
Set the valence for the electrode
Only combination ion electrodes can be used with the D-73 meters
A combination electrode includes a reference sensor
Combination and ion electrodes with a separate reference electrode can be used with the F model meters

## Slide 8: ORP Electrodes
Accurate temperature readings are critical for ORP measurements
ORP standards typically have a short shelf life
Select a standard that is near the expected sample measurement
Standard measurements should be within +/- 15mV. If not replace the Model 300 internal solution.

## Slide 9: Conductivity Electrodes
Input the correct cell constant
Input the electrode cell value
Ensure the temperature of the standard solution is accurate
Meter default settings are ATC, 2% temperature compensation

## Slide 10: Dissolved Oxygen Electrodes
Galvanic electrode method
Set salinity concentration
Calibrate and measure with gentle stirring of the solutions
Perform zero calibration with 50g/L sodium sulfite solution
Perform span calibration in air saturated water